EXPERIMENTAL PROTOCOL Phase 1 & Phase 2 3D Scanning Platform for Tissue Characterization Champalimaud Foundation · Biophotonics Platform

# At a Glance — Which Target Do I Use?

Use this table at the start of every session to confirm which phantom to place on the tray.

| Phase | Test | Target | Why |
| --- | --- | --- | --- |
| Phase 1 | Test 1 — Repeatability | 🔺 Pyramid | Rigid, no movement, known geometry for noise floor |
| Phase 1 | Test 2 — Feedrate sweep | 🔺 Pyramid | Rigid, known geometry — speed effect isolated cleanly |
| Phase 1 | Test 3 — Step size sweep | 🔺 Pyramid | Rigid, known geometry — spacing effect isolated cleanly |
| Phase 1 | Test 4 — Z height sweep | 🔺 Pyramid | Rigid, known geometry — Z effect isolated cleanly |
| Phase 2 | Test A — Geometric verification | 🔺 Pyramid | Known face angles let you verify adaptive-Z accuracy analytically |
| Phase 2 | Test B — Biological relevance | 🧫 Pig skin | Realistic tissue surface — the clinically meaningful result |

Rule of thumb: Pyramid for everything in Phase 1, and for the first part of Phase 2. Only switch to the pig skin for Phase 2 Test B. Do not mix targets within the same test.

⚠ EMERGENCY PROCEDURES — Read Before Starting

Know these before you touch anything. Two levels of stop, use the appropriate one.

| FEED HOLD (pause) | When to use: something looks slightly wrong but there is no immediate collision risk. How: click the Hold button in the GRBL monitor panel. Effect: scanner decelerates smoothly to a stop and holds position. Scan is paused but recoverable. You can resume or abort cleanly from here. |
| --- | --- |

| EMERGENCY STOP (hard reset) | When to use: probe is about to hit the target, unexpected motion, collision risk. How: click Emergency Stop in the GRBL monitor panel. Effect: hard reset sent directly to GRBL, motion stops within milliseconds. Scan is aborted (not recoverable). After emergency stop you MUST re-home before the next scan. After stopping: check the probe is clear of the target before re-homing. Re-home slowly. |
| --- | --- |

Standard Scan Procedure — Use This Every Single Time

Every scan in every test follows these exact steps. Learn them once, repeat them every run.

Step 1 Home the scanner. → GRBL monitor panel → click Home. → Wait until position reads X: 0.000 Y: 0.000 Z: 0.000. → Do not proceed until homing is confirmed.

Step 2 Confirm the camera is live and the ROI is drawn. → The live colour view should be showing the target. → The ROI box should surround the scan area — adjust if needed. → If the ROI has moved since the last run, check that calibration is still valid.

Step 3 Optional: capture a topography preview. → Click the Topography button to see the current depth/height map. → Confirm the target is visible and the peak height looks reasonable. → Note the peak height in mm from the status bar — you will need it for Phase 2.

Step 4 Open the Raster Scan dialog and enter parameters. → Set feedrate, line spacing, Z height, and scan mode as specified for the current test. → Double-check the values before clicking anything.

Step 5 Click Go to ROI Start. → The scanner raises to safe travel Z then moves to the scan start position. → Watch the motion — if anything looks wrong, click Hold immediately. → Wait for the status bar to say 'Reached the prepared raster start at...' → Do NOT click Start Raster until this message appears.

Step 6 Click Start Raster Scan. → G-code streams to GRBL. The scanner moves along the raster pattern. → The overlay on the camera view shows completed lines in real time. → Do not interact with other controls during the scan.

Step 7 Wait for reconstruction. → When the last line finishes, the software runs reconstruction automatically. → Status bar will say 'Streaming reconstruction complete: ...' → Note the run directory path shown — this is where all output files are saved.

Step 8 Record the run in your lab notebook. → Date/time, test name and condition, run directory path, any anomalies.

Step 9 Re-home before the next run. → Always re-home between runs, even within the same test. → This eliminates any accumulated position drift.

PHASE 1 — Geometric and Mechanical Characterization

| 🔺 PYRAMID | Use the pyramid phantom for all of Phase 1. Keep it fixed on the tray throughout. Do not move it between tests or between repeats within a test. |
| --- | --- |

# Phase 1 — Session Setup

Do once per session: Place the pyramid at the centre of the tray. Fix it so it cannot shift (a small piece of tape at the base corners is enough). Perform calibration verification — open the calibration file and confirm xy_homography and plane_model are present. Capture one topography image and record the peak height. This value is your reference for the entire session.

Record in your lab notebook before any scan:

- Date and start time of session

- Calibration file name / save date

- Pyramid peak height from topography (mm)

- Ambient temperature if a thermometer is available

# Phase 1 — Test 1: Repeatability Baseline

## Parameters — set these in the dialog and do not change them for any repeat

| Parameter | Value |
| --- | --- |
| Scan mode | Fixed-Z |
| Feedrate | 800 mm/min |
| Line spacing | 1.0 mm |
| Z height | 10 mm (above tray surface) |

## How many times

8 repeats. Do not touch the pyramid or change any setting between repeats.

## Step-by-step

Follow the Standard Scan Procedure (previous page) for each repeat. Between repeats: re-home only — no other changes.

## What to record per repeat

- Run directory path

- Start time

- Any anomaly (unexpected stop, depth dropout, etc.)

## What to analyse later

- Load all 8 surface_pointcloud.ply files. Fit a plane to each pyramid face in each run.

- Compute standard deviation of fitted plane Z values across the 8 repeats — this is your repeatability σ_Z.

- Target: σ_Z < 0.3 mm. If larger, investigate before continuing.

✓ Checkpoint: Only proceed to Test 2 once you have 8 valid runs saved and σ_Z calculated. Write σ_Z in your notebook — you will reference it in every subsequent test.

# Phase 1 — Test 2: Feedrate Sweep

## Parameters — vary feedrate, keep everything else fixed

| Parameter | Value |
| --- | --- |
| Scan mode | Fixed-Z |
| Line spacing | 1.0 mm |
| Z height | 10 mm |
| Repeats per feedrate | 3 |

## Feedrate values — run in this order

| Run order | Feedrate (mm/min) | Notes |
| --- | --- | --- |
| 1 | 200 | Slowest — baseline spatial quality |
| 2 | 500 |  |
| 3 | 800 | Mid-range — your Test 1 value |
| 4 | 1200 | Fast |
| 5 | 1600 | Very fast |
| 6 | 2000 | Stress test — watch for jitter |

3 repeats at each feedrate = 18 total runs. Re-home before each run.

## What to analyse later

- From motion log: inter-line spacing uniformity (standard deviation of measured Y at line transitions) vs feedrate.

- From reconstruction: RMS residual of reconstructed pyramid face vs known face angle, vs feedrate.

- Plot RMS residual vs feedrate with error bars. Find the elbow — speed above which accuracy degrades.

- The feedrate just below the elbow is your operating feedrate. Write it down — use it for all remaining tests.

✓ Checkpoint: Record your chosen operating feedrate: _______ mm/min. Use this value for Tests 3, 4, and all of Phase 2.

# Phase 1 — Test 3: Line Spacing Sweep

## Parameters — vary line spacing, keep everything else fixed

| Parameter | Value |
| --- | --- |
| Scan mode | Fixed-Z |
| Feedrate | Optimal from Test 2 |
| Z height | 10 mm |
| Repeats per spacing value | 3 |

## Spacing values — run finest to coarsest

| Run order | Line spacing (mm) | Notes |
| --- | --- | --- |
| 1 | 0.5 | Very fine — high quality, slow |
| 2 | 1.0 | Baseline from Tests 1 & 2 |
| 3 | 2.0 | Medium |
| 4 | 3.0 | Coarse |
| 5 | 5.0 | Very coarse — likely undersampled |

3 repeats per value = 15 total runs.

## What to analyse later

- From reconstruction trigger_map.npz: coverage fraction (% of valid pixels) vs spacing.

- RMS face residual vs spacing. Plot both on separate graphs with error bars.

- Choose minimum spacing where coverage > 95% AND residual ≤ 2× σ_Z. This is your operating spacing.

✓ Checkpoint: Record your chosen operating line spacing: _______ mm. Use this for Test 4 and all of Phase 2.

# Phase 1 — Test 4: Z Height Sweep

## Parameters — vary Z height, keep everything else fixed

| Parameter | Value |
| --- | --- |
| Scan mode | Fixed-Z |
| Feedrate | Optimal from Test 2 |
| Line spacing | Optimal from Test 3 |
| Repeats per Z height | 3 |

## Z height values — run largest to smallest

Safety: Always start at the largest Z and work downward. Before running Z = 5 mm, confirm from the topography that the pyramid apex plus standoff leaves clearance — if not, skip that value.

| Run order | Z height above tray (mm) | Notes |
| --- | --- | --- |
| 1 | 20 | Start here — very safe |
| 2 | 15 |  |
| 3 | 12 |  |
| 4 | 10 | Baseline from previous tests |
| 5 | 8 | Close — verify safe before running |
| 6 | 5 | Very close — only run if confirmed safe from topography |

3 repeats per Z value = up to 18 total runs.

## What to analyse later

- RMS face residual vs Z height. Depth noise standard deviation vs Z height.

- Fit a curve to depth noise vs Z to show how noise grows with distance.

- Identify usable Z working range: where residual stays within 2× σ_Z.

- This range directly informs the standoff setting for Phase 2.

✓ Phase 1 complete. You now have: operating feedrate, operating line spacing, and usable Z range. These three values define all scan settings for Phase 2. Write them in your notebook before moving on.

PHASE 2 — Surface-Following Accuracy

Phase 2 compares fixed-Z mode vs adaptive-Z (surface-following) mode on non-flat targets. The central question: does adaptive-Z keep the probe at a more consistent distance from the surface?

All scan parameters (feedrate, line spacing) come from Phase 1. The standoff distance is typically 2.0 mm — adjust if Phase 1 Test 4 indicates a different optimal working distance.

Test A — Pyramid (Geometric Verification)

| 🔺 PYRAMID | Keep the same pyramid from Phase 1. Do not move it. You will run 3 conditions on the pyramid before switching to the pig skin. |
| --- | --- |

# Phase 2 Test A — Session Setup

Before starting Test A, capture a topography image and record:

- Peak height of the pyramid (mm): ___________

- Mean face height (mm): ___________ (read from topography colour scale)

You will use these to set the fixed-Z heights for conditions A1 and A2:

- A1 fixed Z = peak height + 2.0 mm standoff

- A2 fixed Z = mean face height + 2.0 mm standoff

# Phase 2 Test A — Three Conditions

| Condition | Mode | Z setting | Purpose |
| --- | --- | --- | --- |
| A1 | Fixed-Z | Peak height + 2.0 mm standoff | Safe but inefficient — probe far from base |
| A2 | Fixed-Z | Mean height + 2.0 mm standoff | Compromise — some risk at apex |
| A3 | Adaptive-Z | Standoff = 2.0 mm, safety margin = 1.5 mm | Primary result — should be uniform |

3 repeats per condition = 9 total runs. Run A1 first, then A2, then A3.

## Step-by-step for each condition

Follow the Standard Scan Procedure for each run. The only thing that changes between conditions is the Z setting in the dialog.

For condition A3 specifically: In the raster scan dialog, select Surface Following mode and set standoff = 2.0 mm. Confirm in the scan log that multiple different Z heights appear — if all Z values are the same the surface following is not working.

## What to analyse later

- For each run: clearance map = scanner_Z (from motion log) − surface_height (from topography_map.npy).

- Compute mean and standard deviation of clearance (σ_clearance) for each condition.

- Plot clearance distribution histograms for A1, A2, A3 overlaid.

- Plot clearance heatmaps (2D) — A3 should look spatially uniform, A1 and A2 should show a systematic gradient.

- For A1 and A2: compare measured clearance gradient to analytically expected gradient from known face angle. This validates your topography map.

✓ Checkpoint: After 9 runs on the pyramid you are done with all pyramid work. The next step is to switch to the pig skin. Do not re-run anything on the pyramid after the pig skin has been on the tray.

⬇ TARGET SWITCH ⬇ Remove the pyramid. Place the frozen pig skin on the tray. Do NOT proceed to Test B until the sample is properly positioned and a topography image has been captured.

Test B — Frozen Pig Skin (Biological Relevance)

| 🧫 PIG SKIN | Frozen pig skin — thawed on the day of the experiment. Handle with PPE. Pat dry before placing on the tray. Confirm the sample cannot rock or shift. |
| --- | --- |

# Phase 2 Test B — Sample Preparation

Important: Thaw the sample on the day of the experiment only. Do not refreeze after thawing — thermal cycling degrades tissue structure and optical properties.

- Put on PPE (gloves minimum). Follow your institution's tissue handling protocol.

- Thaw the sample. Pat the surface dry with absorbent paper — surface moisture causes specular reflections that degrade depth sensing.

- Place skin-side up on the tray. Ensure it lies flat. If it rocks, use a small amount of modelling clay at the edges (not under the scan area).

- Place a small piece of tape at one corner as a position marker — you will check this between runs to confirm the sample has not shifted.

- Capture a topography image. Record peak height and mean height from the status bar.

- Calculate fixed-Z heights for B1 and B2:

  - B1 fixed Z = peak height + 2.0 mm standoff: ___________

  - B2 fixed Z = mean height + 2.0 mm standoff: ___________

# Phase 2 Test B — Three Conditions

| Condition | Mode | Z setting | Purpose |
| --- | --- | --- | --- |
| B1 | Fixed-Z | Peak height + 2.0 mm standoff | Safe but far from surface at low points |
| B2 | Fixed-Z | Mean height + 2.0 mm standoff | Compromise — contact risk at peaks |
| B3 | Adaptive-Z | Standoff = 2.0 mm, safety margin = 1.5 mm | Primary result — clinically relevant |

3 repeats per condition = 9 total runs. Run B1 first, then B2, then B3.

## Between every repeat

- Check the tape marker — confirm sample has not shifted.

- Re-home the scanner.

- Do not pat or adjust the sample surface between repeats.

## What to analyse later

- Clearance map = scanner_Z (motion log) − surface height (topography_map.npy in run directory).

- Mean and σ_clearance for each condition. Count of near-contact points (clearance < 0.5 mm).

- Clearance heatmap for each condition. Histogram for all three conditions overlaid.

- Main thesis result: σ_clearance ratio between fixed-Z and adaptive-Z on real tissue.

If FLIM is connected: Run condition B3 with pyProbe active. Also record the FLIM intensity map for that run. The CV of intensity across the scan for adaptive-Z vs fixed-Z is your optical result — the mechanical improvement translating into measurement quality.

# Output Files — What the Platform Saves Automatically

For every scan run, the platform creates a timestamped directory containing:

| File | What it contains | Used for |
| --- | --- | --- |
| scan_metadata.json | All scan settings, calibration, Z plan summary | Verify settings after the fact |
| raster_motion_log.csv | Scanner position at 50 Hz with timestamps | Compute clearance maps |
| topography_map.npy | Pre-scan surface height in mm (calibrated) | Reference surface for clearance calculation |
| topography_map.png | Colour-mapped height image with scale bar | Figures for thesis |
| reconstruction/surface_pointcloud.ply | 3D point cloud of the tissue surface (XYZ mm, coloured by height) | Geometric analysis, 3D figures |
| reconstruction/trigger_map.npz | 2D coverage map: trigger index and height per pixel | Coverage fraction calculation |
| reconstruction/surface_topography.png | Height map with scan path overlay | Figures — shows what was scanned |

# Lab Notebook Entry — Record for Every Run

| Field | Notes |
| --- | --- |
| Date / time |  |
| Phase and Test | e.g. Phase 1 Test 2 — feedrate 1200 mm/min, repeat 2 of 3 |
| Target | Pyramid / Pig skin |
| Feedrate (mm/min) |  |
| Line spacing (mm) |  |
| Z height / mode | Fixed-Z at ___ mm OR Adaptive-Z, standoff ___ mm |
| Run directory path |  |
| Topography peak height | ___ mm (from topography capture before this run) |
| Any anomalies | Emergency stop? Depth dropout? Unexpected motion? |
| Ambient temperature | ___ °C (if thermometer available) |